| 1. Deviation Report/偏差汇报 |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Deviation Title 偏差标题 |  | Unexpected solid residue and stain was found in the glove box of solid-phase peptide synthesizer |  |  |  |  |  |  |
| Dept./部门 |  | Production Dept. |  | Project No./项目编号 |  |  |  | N/A |
| Step/Batch/步/批 |  | N/A |  | Ins No./操作指令号 |  |  |  | N/A |
| Date happened/ 发生日期 |  | 10/10/2026&10/15/2026 |  | Date discovered/ 发现日期 |  |  |  | 10/16/2026 |
| Project Category 项目类别 |  | ☐ Category 1/类别一项目 ☐ Category 2/类别二项目 ☐ Category 3/类别三项目 ☐ Category 4/类别四项目 ☒ Non-direct project related/非直接与项目相关 |  |  |  |  |  |  |
| Deviation Category 偏差类别 |  | ☐ Critical deviation/关键偏差 ☐ Major deviation/主要偏差 ☒ Minor deviation/次要偏差 |  |  |  |  |  |  |
| Deviation Attribution/ 偏差归属 |  | ☐Process parameter out/工艺参数超标 |  |  | ☐Yield out/收率超标 |  |  |  |
|  |  | ☐Power /utility /动力/公用工程 |  |  | ☐Equipment qualification/设备确认 |  |  |  |
|  |  | ☐Environment control/环境控制 |  |  | ☒Cleaning and disinfection/清洁和消毒 |  |  |  |
|  |  | ☐Raw data management/原始数据管理 |  |  | ☐Computerized system control/计算机系统控制 |  |  |  |
|  |  | ☐Supplier control/供应商控制 |  |  | ☐Calibration management/校准管理 |  |  |  |
|  |  | ☐Document control/文档控制 |  |  | ☐Labeling/标签 |  |  |  |
|  |  | ☐Extraneous matter/异物 |  |  | ☐Other/其他 N/A |  |  |  |
| Deviation description/ 偏差内容 |  | On 10/16/2026 at 11:50, during on-site visit in Workshop 13, Building 12 at Asymchem Life Science (Tianjin) Co., Ltd. (Address: 265 South Street, Binhai New Area, TEDA, Tianjin, China), hereafter referred to as the TJ4 site, Lilly customer representatives discovered two unexpected cases, while inspecting the interior of glove box fixed on solid-phase peptide synthesizer TJ4S-1213-Sy15 and TJ4S-1213-Sy16. Two cases are detailed as follow: Case 1: Solid residue at the outer surface of discharging outlet, in the glove box of solid-phase peptide synthesizer TJ4S-1213-Sy15 (refer to Pic. 1) is found; Case 2: Stain at the corner of ziplock LDPE bag which is used to protect the sampler TJ4S-1213-Sy16-Sam01 (refer to Pic. 2) (the sampler is placed in the glove box of solid-phase peptide synthesizer TJ4S-1213-Sy16) is found. Pic. 1 Solid residue at the outer surface of discharging outlet; Pic. 2 Stain at the corner of PE bag which is used to protect the sampler. Although both synthesizers are at the status of “Cleaning” after passivation, but both cases are unexpected, hence this deviation was initiated. |
|  |  | (Information such as number, date, time, location and incident, the name of any persons involved, witnessing the incident in addition to the observer and so on shall be clearly stated, as well as departure from which instruction, standard or procedure, etc./清晰简明地阐述事件相关的信息，如编号、日期、时间、地点、事件、发现人、任何有关的人员及事件的证明人，及所偏离的指令、标准、规程) |  |  |  |  |  |  |
| Initial Impact Assessment / 初始风险评估 |  | Synthesizer TJ4S-1213-Sy15 and TJ4S-1213-Sy16 are still under cleaning/TOC testing stage. They are not used for next product production (Per Asymchem procedure, after the cleaning and testing, there is final inspection and QA release before use in next product), this residual resin has been cleaned/removed after the deviation occurring, and LDPE packing bag for sampler has been replaced to a new one. Both of those two synthesizer shall not be used until they are further cleaned and released. After initial impact assessment, it has no impact on product quality. (Initial assess the impact extend and scope to the system/equipment/product/procedure/method. /对于系统，设备/产品、分析方法、规程的影响范围和程度的评估) |  |  |  |  |  |  |
| Emergency Action/ 应急措施 |  | Since the equipment was in the process of cleaning/TOC testing stage, the status label of the equipment was under “Cleaning”, hence no emergency action was taken. |  |  |  |  |  |  |
|  |  | If no human error is found at this stage, but found as the investigation goes on, the training can be provided at the end of this investigation in the preventive action/如果当前未发现人员错误，但是随着调查的进行发现人员错误，相关培训可以在本调查结束后作为预防措施执行 ☐ No Human error is found/未发现人员错误 If any personal error is found, the direct personnel involved in the event shall be trained /如发现有人员错误，应对事件直接当事人进行相关培训 ☐ Training is required and finished, can continually perform quality related works. / 需要培训，已完成培训，当事人可继续质量相关工作。 ☒ Training is required and to be trained prior to performing quality related works. / 需要培训待进行培训，接受培训前不得参与质量相关工作。 Attached supporting document /支持性附件: Deviation-TJ420260226.01 Attachment 01 Wider training scope (if there is) can be provided at the end of this investigation in the preventive action / 更大范围的培训可以在本调查结束后作为预防措施执行 |  |  |  |  |  |  |
| Filled by/Date/ 填写人/日期 |  |  | Dept. Reviewer/Date 部门复核人/日期 |  |  |  |  |  |
| Input by QA/ QA填写 | QAA with client/客户质量协议 |  |  |  |  | ☒ Yes/有 ☐ No/无 |  |  |
|  | Notify customer/通知客户 |  |  |  |  | ☒ Yes/是 ☐ No/否 ☐No need/不需要 |  |  |
|  | Customer approval/是否客户审批 If “Yes”, <Client Approval Sign Page> (QF01086) should be sent to client for approval if necessary, and attached. /如果“是”，必要时向客户发送《客户审批签字页》（QF01086），并将客户签字后的审批页附后。 |  |  |  |  | ☐ Yes/是 ☒ No/否 ☐ No need/不需要 |  |  |
|  | QA Reviewer /Date/QA复核人/日期 |  |  |  |  |  |  |  |

| 2. Deviation investigation and cause analysis /偏差调查及原因分析 |  |
| --- | --- |
| 2.1 Personnel related/人员因素 |  |
| A review was conducted of all work activities and personnel operations involved in solid-phase peptide synthesizers TJ4S-1213-Sy15 and TJ4S-1213-Sy16 along with the whole equipment use-log, to investigate the potential source of solid residue in the glovebox and stain on the protective bag. Case 1 for Solid residue: The equipment use logbook has been reviewed, the historical usage of solid-phase peptide synthesizer TJ4S-1213-Sy15 is listed in Table 1: Table 1 Historical Usage of solid-phase peptide synthesizer TJ4S-1213-Sy15 Usage; Start time and end time. Initial Equipment Qualification and Trial Run; 02/07/2026~03/25/2026. Cleaning after Qualification and Trial Run; 03/27/2026~03/31/2026. Solvent Run before Full Scale Production; 04/27/2026~04/29/2026. Cleaning before Full Scale Production; 05/03/2026~05/08/2026. Production for Full Scale Batch CPo140020-03-01-01; 05/20/2026~06/02/2026. Cleaning after Full Scale Production; 06/02/2026~06/08/2026. Supplement Equipment Qualification; 06/08/2026~08/18/2026. Cleaning and Chemical Residue Test after Qualification; 08/20/2026~08/27/2026. Supplement Equipment Qualification and Solvent Run; 08/28/2026~09/17/2026. Pipeline Passivation; 09/17/2026~10/08/2026. Cleaning after Pipeline Passivation; Starting from 10/13/2026, swab sample for TOC was taken on 10/15/2026 from the discharging port in glove box. The Solid-Phase Peptide Synthesizer TJ4S-1213-Sy15 was inspected and released by QA on 08/22/2026 (the latest QA inspection time) with no solid residue identified. Between the release date and the solid residue discovery date on 10/16/2026, Solid-Phase Peptide Synthesizer TJ4S-1213-Sy15 has been used six times, including: - Supplemental qualification from 08/28/2026 to 08/29/2026 (software optimization only, no material feeding) - Solvent run from 08/30/2026 to 09/01/2026 (Lot No.: CPo140020-03-01Test-02 for solvent run with DMF and Piperidine) - Supplemental qualification from 09/02/2026 to 09/15/2026 (software optimization only, no material feeding) - Solvent run from 09/16/2026 to 09/17/2026 (Lot No.: CPo140020-03-01Test-03 for solvent run with DMF and Piperidine) - Pipeline passivation from 09/17/2026 to 09/30/2026 (NaOH, citric acid, EDTA solution in purified water used for passivation) - Equipment train cleaning from 10/13/2026, and equipment train cleaning is on-going now. (Purified water and DMF). The last operation is the sampling for TOC test on 10/15/2026. Based on the synthesizer design, the discharge port within the glovebox is designed to collect peptide-on-resin after synthesis. Liquid streams such as solvents and solutions used in above six running from 1 through 6 shall be discharged via the discharge pipeline. It is suspected that the solid found in the bottom of glove box is peptide-on-resin from Full Scale Batch CPo140021-03-01-01. To confirm this hypothesis, IR identity (ID) analysis is performed against the manufactured peptide-on-resin reference sample. According to the analysis result (refer to Deviation-TJ420260226.01 Attachment 02), the solid residue was confirmed to be peptide-on-resin from batch CPo140020-03-01-01. Since the Solid-Phase Peptide Synthesizer TJ4S-1213-Sy15 had already been cleaned and released, the equipment was disassembled at the glovebox discharge port to identify the source of the peptide-on-resin. Following disassembly, a hard-to-clean area was observed at the bottom of bolt used to fix the blind plate on the discharge port. See Table 2 below for the detailed location: Table 2 Step by step Zoom-in inspection of the Location of residual peptide-on-resin Fig 1, Overall appearance; Fig. 2, Detail front appearance; Fig 3, Further zoom-in appearance. After disassembly of the blind plate on the discharge port, no visible peptide-on-resin residue was observed either inside or outside the discharge port.; From a routine check on the bolt that fixes the blind plate on the discharge port, no obvious peptide-on-resin residue was observed from the front view.; When further zoom in the bottom of bolt under flashlight, trace power (similar with the observed powder) was found adhered to the rear side of the screw. On 10/14/2026, swab sample was taken through the discharging port for TOC testing. During sampling, the blind plate was disassembled, it was suspected that trace peptide-on-resin adhering at the back of blot fall down to the glove box due to vibration during blind plate disassembly and assembly. Then solid was found on 10/15/2026.. From the above investigation, it can be seen that trace peptide-on-resin residue adhered to bottom of bolt equipment release was not fully removed in the equipment cleaning and not identified in the equipment release in June and August 2026. Therefore, human factor (fail to ensure the cleanliness of equipment) is the direct cause of this deviation. Further investigation on root cause of Case 1 utilizing Human Error Analysis Toolkit and 5 WHY investigation methods on the residual peptide-on-resin is performed in Section 2.3 Method Review. Case 2 for Stain: The equipment use logbook has been reviewed, the historical usage of solid-phase peptide synthesizer TJ4S-1213-Sy16 is listed in Table 3: Table 3 Historical Usage of solid-phase peptide synthesizer TJ4S-1213-Sy16 Usage; Start time and end time. Initial Equipment Qualification and Trial Run; 02/24/2026~04/24/2026. Cleaning after Qualification and Trial Run; 04/25/2026~04/27/2026. Solvent Run before Full Scale Production; 05/04/2026~05/05/2026. Cleaning before Full Scale Production; 05/06/2026~05/13/2026. Production for Full Scale Batch CPo140021-03-01-01; 05/18/2026~05/28/2026. Cleaning after Full Scale Production; 05/30/2026~06/02/2026. Supplement Equipment Qualification; 06/04/2026~08/18/2026. Cleaning and Chemical Residue Test after Qualification; 08/21/2026~08/24/2026. Supplement Equipment Qualification and Solvent Run; 08/26/2026~09/16/2026. Pipeline Passivation; 09/16/2026~10/11/2026. Calibration of Jacket Temperature Transmitter; 10/08/2026~10/10/2026 2nd-shift from 08:00. Tightness Check before Cleaning; 10/10/2026 3rd-shift from 16:00. Cleaning after Pipeline Passivation; Starting from 10/11/2026. Review of the equipment usage logbook confirmed that the sampling tool was cleaned on 06/02/2026 after sampling during full-scale production of batch CPo140021-03-01-01. The cleaned sampling tool was placed in a new LDPE ziplock bag for storage, and QA affixed the “Accessory Equipment Release Label”. Between release date and the stain discovery date on 10/16/2026, Solvents were involved on the Solid-Phase Peptide Synthesizer TJ4S-1213-Sy16, including: - Ethylene glycol was used during qualification at the period of 06/04/2026~07/25/2026 - DMF&MTBE were used during cleaning at the period of 07/30/2026~09/05/2026 - Citric acid & NaOH & EDTA solution in purified water was used during passivation at the period of 09/16/2026~09/30/2026 IR analysis was performed to confirm the chemical identity of the stain observed at the corner of the LDPE ziplock bag. According to the analysis result (refer to Deviation-TJ420260226.01 Attachment 03), the stain was identified by IR as ethylene glycol. Ethylene glycol is used as a solvent during equipment qualification (but Ethylene glycol used for qualification is colorless, not pale yellow), also serves as the heating and cooling medium within the TCU (The color of Ethylene glycol in TCU medium is pale yellow). After review the historical usage records of solid-phase peptide synthesizer TJ4S-1213-Sy16, latest three operations are: - Cleaning after passivation, where ethylene glycol isn’t involved; - Airtightness check before Cleaning by Yugang Li, where ethylene glycol isn’t involved; - Calibration of Jacket Temperature Transmitter by Xinjie Zhang, where ethylene glycol isn’t involved. While considering the jacket medium is ethylene glycol, the detail progress during calibration of jacket temperature transmitter is further reviewed. On 10/08/2026, temperature transmitters CYJ11841 and CYJ11828 (installed to monitor TCU jacket temperature) were removed for calibration. On 10/10/2026 2nd shift from 08:00, Operator Xinjie Zhang carried out reinstallation of all calibrated transmitters. Interviewing with the Operator Xinjie Zhang, no abnormality found during reinstallation. While further interviewing with the Operator Yugang Li (the operator of 3rd shift from 16:00) who is responsible for airtightness check of synthesizer TJ4S-1213-Sy16 on 10/10/2026, small amount of ethylene glycol was found around the point of TCU temperature transmitter CYJ11828, then cleaned the ethylene glycol; however, trace ethylene glycol might be inadvertently adhered to his gloves. Then airtightness check was further performed by Operator Yugang Li, in order to ensure the airtightness, Yugang Li opened the glove box to inspect the discharge port tightness. The LDPE ziplock bag containing the sampling tool was located nearly the discharge port, so the Yugang Li moved LDPE ziplock bag to another position within the glove box. It is speculated that trace ethylene glycol on the operator’s gloves came into contact with the LDPE ziplock bag, resulting in the stain observed on the LDPE bag. The ethylene glycol medium inside the jacket is pale yellow, exhibiting a color similar to that of the stain found on the LDPE ziplock bag. Procedure defect is the root cause. There is no detailed glove change requirement when dealing with non-process potential exposed operation. In addition, it is not appropriate to put release label on the sampler since the synthesizer is under cleaning and TOC testing. For further investigation, refers to 2.3 method factor section. (Also interview of the people involved in doing the work where the incident occurred/面谈偏差现场执行相关操作的人员） |
| 2.2 Equipment/设备因素 |  |
| After investigation, the equipment solid-phase peptide synthesizers TJ4S-1213-Sy15 and TJ4S-1213-Sy16 has been successfully qualified. Their qualified functions are suitable for production use and cleaning. The observed powder and stain was not rusted metal, therefore, no equipment-related factors were identified as the root cause of this deviation. (Includes any emergency maintenance review/包括任何紧急维护的回顾) |  |
| 2.3 Method Review/方法因素 |  |
| Case 1 for Solid residue: Further investigation utilizing Human Error Analysis Toolkit (QM01026, refer to Table 3) and 5 WHY investigation methods (refer to Table 4) on the residual peptide-on-resin is performed as follow: Table 3. Root cause analysis on equipment cleaning and cleanliness inspection – Human Error Analysis Toolkit Factor; Verification Question; Verification Result; Justification. Yes; No; N/A. 1. Error type classification. Skill-based mistake (slip or lapse); The operator is familiar with the task, and the task is routine, and performed almost automatically.; ☒; ☐; ☐; The operator responsible for equipment cleaning worked on equipment cleaning for many times following equipment cleaning method, cleaning manual and cleaning SOP. The related personnel worked on equipment release for many times following equipment cleaning method and related SOP for equipment cleaning and release based on the defined cleaning checklist, the related personnel is familiar with the relevant document. As conclusion, there’s missing information regard with worst case cleaning point in the risk assessment for cleaning of peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01 refer to Deviation-TJ420260026.01 Attachment 04), detail refer to the 5 WHY investigation methods (refer to Table 4) in Section 2.3 Method Review.. Rule-based mistake; The person intended to follow a rule/procedure but applied the wrong one (e.g., wrong SOP version).; ☐; ☒; ☐; Only latest version cleaning method and checklist for equipment cleaning is used during equipment cleaning verification by QA.. Knowledge-based mistake; The task is novel, rare, has no clear procedure, or requires problem-solving.; ☐; ☒; ☐; The related personnel has been trained and qualified for equipment cleaning and equipment release, and related personnel work on equipment cleaning and cleanliness inspection for many times, they were familiar with the relevant document.. Intentional violation mistake; The person knew the correct procedure but deliberately did not follow it (shortcut, convenience).; ☐; ☒; ☐; QA is the qualified personnel for equipment release, with no intentional violation during equipment release.. 2.Performance influencing Factors. Fatigue; Is the person working overtime or consecutive night shifts?; ☐; ☒; ☐; The equipment cleaning and release were completed in regular working hours, and no overtime or consecutive night shifts were involved.. Time pressure; Was there a production/release deadline (month-end, shipment, etc.)? Was the person asked to “speed up”?; ☐; ☒; ☐; Sufficient time was available for equipment cleaning and release with no production/release deadline requirement.. Interruption distraction; Was the person interrupted by phone call, colleague question, alarm during the critical step?; ☐; ☒; ☐; During the equipment cleaning and release, no person interrupted was involved.. Multi-tasking; Was the person required to perform two or more attention-demanding tasks simultaneously?; ☐; ☒; ☐; The related personnel only performed one attention-demanding task at one time when doing the relevant works.. Memory load; Did the task require memorizing >5-7 independent steps or numbers without any written aid?; ☐; ☒; ☐; The equipment cleaning logbook has specified the key point of equipment cleaning and release, with no tough memory load.. Emotional/psychological stress; Was there any family, health, or work-related pressure known at the time?; ☐; ☐; ☒; During the equipment and cleaning, no family, health, or work-related pressure to relevant person were involved.. 3.Procedure and Documentation Factors. Readability; Is the SOP written in long paragraphs without flowcharts, pictures, or tables?; ☐; ☒; ☐; There is approved cleaning logbook as the checklist to perform equipment cleaning and cleanliness inspection, hence the written SOP is readable.. Completeness; Is any critical step missing? Does it rely on “common sense” or “experience”?; ☒; ☐; ☐; For the cleaning record, there is requirement to clean and check bolt and nuts, etc. However, it lacks of detailed instruction on cleaning and check on the worst case, bottom of bolt at the discharging port.. Version control; Is the person using the current effective version?; ☒; ☐; ☐; The equipment cleaning and cleanliness inspection were performed based on latest effective version cleaning logbook and cleaning method.. Consistency; Are terminology and format consistent across related SOPs?; ☒; ☐; ☐; The description in cleaning method and cleaning logbook were consistent of SOP for cleaning with no misunderstanding opportunity.. Accessibility; Can the person easily access the SOP at the work location (e.g, without leaving sterile area)?; ☒; ☐; ☐; The cleaning method and cleaning manual has been distributed to the production are for person’s easy access, and the cleaning logbook is on-site for person’s filling.. Alignment with practice; Does the procedure match actual on-site operations? Are there any undocumented or informal practices?; ☐; ☒; ☐; The cleaning method, cleaning manual and cleaning logbook match the actual on-site operation and cleanliness inspection.. 4.Training and Competence Factors. Initial training completed; Has the person received formal training on this task with records?; ☒; ☐; ☐; The relevant personnel in production department did receive equipment cleaning related training prior to equipment cleaning. The QA did receive equipment release related training prior to equipment cleanliness inspection.. Training content adequacy; Does training include hands-on practice (not only classroom)? Does it include “why” (rationale) and abnormal handling?; ☒; ☐; ☐; Training related to equipment cleaning and cleanliness inspection includes the hands-on practice with “why” and abnormal handling.. Supervised practice; Is there a sufficient period of supervised practice to verify competency post-training?; ☒; ☐; ☐; After receiving relevant training, the personnel concerned had performed numerous similar tasks and possessed the skills required to fulfill the job responsibilities.. Refresher for infrequent tasks; Is there a periodic review or skill maintenance plan for tasks not performed regularly?; ☐; ☒; ☐; The personnel concerned had performed numerous similar tasks and possessed the skills required to fulfill the job responsibilities.. Knowledge retention; Can the person correctly recall key steps and risks?; ☒; ☐; ☐; The personnel concerned had performed numerous similar tasks and well knew the key points in equipment cleaning and cleanliness inspection.. 5.Tools, Equipment and Environment Factors. Similarity confusion; Are different materials, containers, buttons, or connectors similar in shape, color, or location?; ☒; ☐; ☐; The peptide-on-resin adhere to rear side of screw is not similar in shape and color, but the location is hard to check. Further review the risk assessment for cleaning of peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01), this hard-to-clean location is missed. Ergonomics; Does the operation require excessive reaching, bending, force, or is it prone to miss-touch?; ☒; ☐; ☐; The cleaning and inspection of the location need to get closer to the screw, which is hard-to-clean and hard-to-check location.. Lighting & noise; Is the lighting in the working area sufficient? Are there any distracting noise or vibration?; ☐; ☒; ☐; Utilizing flashlight to check the cleanliness, hence the lighting is sufficient during cleanliness inspection.. Labeling; Are pipes, valves, switches clearly, durably, and logically labeled?; ☒; ☐; ☐; The pipes, valves, switches of equipment are clearly, durably, and logically labeled.. Error-proofing; Is there any physical poka-yoke (e.g., different connector sizes) or forced function? If not, can it be added?; ☐; ☒; ☐; There’s no physical poka-yoke to avoid error currently, since this hard-to-clean location is missed in the risk assessment for cleaning of peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01). 6.Management and Culture Factors. Informal rules; Is there an unofficial working method “everyone does it this way” practice?; ☐; ☒; ☐; During the equipment cleaning and cleanliness inspection, it is hard to clean location, since this hard-to-clean location is missed in the risk assessment for cleaning of peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01), hence no unofficial working method “everyone does it this way” practice was involved.. Unreasonable procedures; Are rules impractical, forcing people to find workarounds?; ☐; ☒; ☐; The equipment cleaning and cleanliness inspection is impractical, but fail to catch the hard-to-clean location in the risk assessment for cleaning of peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01). Supervision effectiveness; Does the supervisor regularly observe on-site operations? Were issues observed but not corrected or escalated?; ☒; ☐; ☐; The supervisor regularly observe on-site operations, the operations follow the defined cleaning method, cleaning manual and cleanliness inspection SOP, no similar issue observed.. Consequence awareness; Does the person understand the quality/safety consequences of not following procedures?; ☒; ☐; ☐; The relevant personnel in production department did receive equipment cleaning related training prior to equipment cleaning. The QA did receive equipment release related training prior to equipment cleanliness inspection. Both well understand the quality consequences of not following procedures.. Blame culture; Is the person afraid of being punished for reporting errors?; ☐; ☒; ☐; Asymchem did not operate under a blame-oriented culture, and the relevant personnel had no concerns when equipment cleaning and cleanliness inspection.. Resource adequacy; Is the person asked to work fast without sufficient manpower, time, or tools?; ☐; ☒; ☐; Sufficient time was available for equipment cleaning and release with no production/release deadline requirement.. Management oversight; Does management regularly review human error trends and drive systemic improvements?; ☒; ☐; ☐; The management regularly review human error trends, but it is hard to clean location, and this is first time to find this location, systemic corrective actions will be taken to avoid future error.. The analysis results with Human Error Analysis Toolkit were presented in Table 3. As investigation above, the error type was classified as “Skill-based mistake (slip or lapse)”, and more analysis were performed across multiple dimensions, including performance-influencing factors, procedure and documentation factors, training and competency factors, tools, equipment and environmental factors, and management and culture factors. The analysis indicated that the bolt back location was omitted in the risk assessment for cleaning of the peptide solid-phase synthesizer at the TJ4 site (RA-TJ42026048.01). The inadequate corresponding guidance documents resulted in failure of production personnel to clean this location, and QA also failed to perform cleanliness inspection on this location, because the QA cleaning checklist contained inadequate instructions and did not include requirements for checking the worst-case contamination point. This was summarized in Table 4 via five-why method. Table 4. 5-Why analysis for Solid residue Why-#; Question; Answer. Why-1; WHY the peptide-on-resin was found?; On 10/14/2026, swab sample was taken from the discharging port. During sampling, the blind plate need to be disassembled and then assembled after TOC sample are pulled, trace peptide-on-resin adhering at the back of bolt falls off due to vibration.. Why-2; WHY the trace peptide-on-resin was stick at the back of screw?; The production operator responsible for equipment cleaning failed to clean the area at the back of the bolt. The QA inspector performing the cleanliness check also did not find the trace peptide-on-resin residue adhering to the back of the bolt because the bolt was not separately screwingly disassembled.. Why-3; WHY the trace peptide-on-resin was not cleaned during equipment cleaning?; The production operator responsible for equipment cleaning follows the cleaning method, cleaning manual and cleaning SOP to perform equipment cleaning. However, the detailed cleaning instruction to worst case area, bottom of bolt at the discharging port, was not included in the cleaning record. Consequently, the production operator failed to clean this location containing trace peptide-on-resin residue.. Why-4; WHY the trace peptide-on-resin was not found during cleanliness inspection?; QA also followed the cleaning record checklist to verify that the equipment (including all components) had been cleaned. As mentioned above, due to detailed instruction inadequate in cleaning checklist, Consequently, the QA failed to check this location containing trace peptide-on-resin residue.. Based on this analysis in Table 4, it can be concluded that although personnel responsible for equipment cleaning and cleanliness inspection have extensive relevant work experience, they performed their work solely in accordance with the existing SOP, risk assessment and related documents. The area at the back of the bolt was not included in the risk assessment and cleaning record/cleaning checklist. Consequently, the production operator failed to clean this location where trace peptide-on-resin residue, and QA also failed to inspect this location. Therefore, the inadequate risk assessment and cleaning record/cleaning checklist was identified as the root cause of Case 1 in this deviation. Case 2 for Stain: According to the hypotheses and analysis presented in Section 2.1, the inadvertent introduction of trace ethylene glycol into the glove box was the direct cause. Further review of SOP DS-MM-015 <Processing Aid Materials Management Procedure (DS)> shows that the SOP clearly defines the required frequency for glove changes during production activities, but it does not specify glove management requirements for non-process operations. Therefore, procedure deficiency is concluded as the root cause of introducing ethylene glycol. In addition, after equipment cleaning, multiple activities were completed. The equipment status of solid-phase peptide synthesizer TJ4S-1213-Sy16, together with all its accessories (including sampling tools and rake placed inside the glove box), should have been updated. However, the sampling tools with “Accessory Equipment Release Label” were inconsistent with the actual status of solid-phase peptide synthesizer TJ4S-1213-Sy16. Further review on SOP for equipment cleaning DS-GMP-003 <Equipment cleaning management procedure for GMP production (DS)>, it showed that the procedure clearly requires status labels for key equipment to be updated, and fails to provide detailed instructions on status label changes for dedicated equipment accessories. Therefore, the procedure deficiency was identified as the root cause of Case 2 in this deviation. (Includes but is not limited to the process, analytical method, standard procedure or operation guideline and so on./ 包括但不限于工艺，分析方法，相关的规程或操作规范等) |
| 2.4 Material Review/物料因素 |  |
| After investigation above, this deviation is not related to material factors. (Includes supply chain and supplier/包括供应链和供应商) |  |
| 2.5 Environment Review/环境因素 |  |
| The investigation confirmed that all production operations involving solid-phase peptide synthesizers TJ4S-1213-Sy15 and TJ4S-1213-Sy16 were performed within the GMP workshop environment. No abnormal operating conditions occurred during this period. Therefore, environmental factors can be ruled out. |  |
| 2.6 Others /其他因素 |  |
| After investigation, there’s no other factors. |  |
| 2.7 Supplementary sheet to Deviation Report/偏差文件附页 |  |
| N/A |  |
| Attached information was Provided by/Date 附页信息提供人/日期 |  |

| 2.8 Root Cause Attribution/根本原因归属 |  |
| --- | --- |
| ☐ Production operation related/生产操作相关 | ☐ Material related/物料相关 |
| ☐ Production equipment failure/生产设备故障 | ☐ Utility failure/公用工程故障 |
| ☒ Procedure defect/规程缺陷 | ☐ Testing operation related/检测操作相关 |
| ☐ Testing method related/分析方法相关 | ☐ Lab equipment failure/分析设备故障 |
| ☐ Environmental control failure/环境控制缺陷 | ☐ Unconscious error/无意识的错误 |
| ☐ Process related/工艺相关 | ☐ Other/其他 N/A |
| 2.9 Root Cause Conclusion /根本原因结论 |  |
| Case 1 for Solid residue: Procedure deficiency is the root cause. The back/bottom of the bolt at discharging port of synthesizer was not included in the risk assessment and cleaning record/cleaning checklist as one of the hard-to-clean locations. Human error is the direct cause. The production operator failed to clean this location where trace peptide-on-resin residue was found, and QA also failed to check this location by require screwing disassemble for inspection from different directions. Case 2 for Stain: Procedure deficiency is the root cause. SOP DS-MM-015 <Processing Aid Materials Management Procedure (DS)> clearly defines the required frequency for glove changes during production activities, but fails to specify glove management requirements for non-production routine operations. The sampler with “Accessory Equipment Release Label” were inconsistent with the actual status of solid-phase peptide synthesizer TJ4S-1213-Sy16. And SOP DS-GMP-003 <Equipment cleaning management procedure for GMP production (DS)> does not provide detailed instructions on status label changes for dedicated equipment accessories. |  |

| 3. Impact assessment/影响评估 |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| 3.1 Deviation impact on quality management system/偏差对质量管理体系的影响 |  |  |  |  |  |  |
| Code 代号 | Deviation Impact Checklist/ 偏差影响检查清单 | Current Doc. No./当前文件编号 | Task Prior to 1st impacted batch release/受影响第一批放行前需完成工作 | Post release Task/放行后需要完成工作 | Responsible person/ 落实人 | Due Date/ 落实时间 |
| A | ☐ Processing techniques /process flow chart update/工艺/工艺流程图更新 | N/A | ☐ | ☐ | N/A | N/A |
| B | ☐ MOR/OR/packaging operation record update/ MOR/OR/批包装记录更新 | N/A | ☐ | ☐ | N/A | N/A |
| C | Validation /验证： ☐ Process validation/工艺验证 ☐ Cleaning validation/清洁验证 ☐ Computerized system validation/计算机系统验证 ☐ IQ,OQ,PQ(system/Equipment name/no.)/IQ,OQ,PQ(系统/设备名称/编号： N/A ) | N/A | ☐ | ☐ | N/A | N/A |
| D | ☐ Method update/ verification/ validation/ transfer/方法更新/确认/验证/转移 | N/A | ☐ | ☐ | N/A | N/A |
| E | Spec. update/质量标准更新： ☐ RM/原辅料 ☐ Packaging material/包装材料 ☐ Intermediate/中间体(中间产品) ☐ Final product/产品 | N/A | ☐ | ☐ | N/A | N/A |
| F | ☐ Stability/稳定性 | N/A | ☐ | ☐ | N/A | N/A |
| G | ☐ Homogeneity study/均一性研究 | N/A | ☐ | ☐ | N/A | N/A |
| H | Document update/文件更新： ☐ SMF ☐ VMP ☐ CVMP ☐ CSVMP ☐ (E)DMF ☐NDA/ANDA ☐ COAs ☒ SOP ☐ QM ☒ QF ☐ Quality Agreement/质量协议 ☐ Analytical Record Template/分析记录模板 ☐ Master Label/主标签 ☒ Other/其他： Risk Assessment for Cleaning of Peptide Solid-Phase Synthesizers at the TJ4 Site | DS-GMP-003; DS-MM-015; QF03376.01; RA-TJ42026048.01; | ☐ | ☒ | Xiongxiong Anyao Tang | 10/25/2026 |
| I | ☐ PI&D update/管道仪表布置图更新 | N/A | ☐ | ☐ | N/A | N/A |
| J | QAA update /QAA更新：☐ Contractor/委托方 ☐ Supplier/供应商 ☐ Client/客户 ☐ Between Asymchem sites/Asymchem厂区间 ☐ Regulatory Authorities/官方机构 N/A | N/A | ☐ | ☐ | N/A | N/A |
| K | Notification letter/通知信件: ☐ Client/客户 ☐ Supplier/供应商 ☐ Regulatory Authorities/官方机构 N/A | N/A | ☐ | ☐ | N/A | N/A |
| L | ☐ Design/Facilities/设计/厂房： —Area to be change/待更新的区域： N/A | N/A | ☐ | ☐ | N/A | N/A |
| M | ☐ Other /其它： N/A | N/A | ☐ | ☐ | N/A | N/A |

3.2 Summary of the impact Assessment/影响评估小结 - The impact on the subsequent batches: Synthesizer TJ4S-1213-Sy15 and TJ4S-1213-Sy16 are still under cleaning/TOC testing stage. They are not used for next product production (Per Asymchem procedure, after the cleaning and testing, there is final inspection and QA release before use in next product), this residual resin has been cleaned/removed after the deviation occurring, it has no impact on next product quality; Sampler TJ4S-1213-Sy16-Sam01 is dedicated accessory to Synthesizer TJ4S-1213-Sy16. The sampler shall be used, cleaned and released together with Synthesizer TJ4S-1213-Sy16. At the time of deviation occurrence, Synthesizer TJ4S-1213-Sy16 was still under cleaning and TOC testing, not released. After TOC testing, the equipment, including Synthesizer TJ4S-1213-Sy16, shall be disassembled cleaned, then released. This sampler packing contamination has no impact on the following batches. However, it is not appropriate to put release label on the sampler since the synthesizer is under cleaning and TOC testing. The SOP DS-GMP-003 <Equipment cleaning management procedure for GMP production (DS)> shall be updated to give a more detailed instruction on the equipment dedicated accessory status label change. - The impact on the previous batches: After Synthesizer TJ4S-1213-Sy15 and TJ4S-1213-Sy16 were qualified and released, they were only used for LARRA Fragment 2 and Fragment 3 full scale batch. That means before Fragment 2 and Fragment 3 full scale batch, those two pieces of equipment were new equipment, and not used for any other product. So this deviation has no impact on previous Fragment 2 and Fragment 3 full scale batch. - The impact on PPQ, cleaning validation and stability/re-test period: The PPQ and cleaning validation have not been commenced. A series of CAPAs shall be made and implemented to address the gap prior to PPQ and cleaning validation. This deviation has no impact on the PPQ and cleaning validation. This deviation is regarding with equipment cleaning and contamination, it has no impact on the stability and re-test period. (Include impact to this batch, other batch of this product, and other product；as well as the validation status and filing impact/包括对于本批次、本产品其的其他批次，以及对于其他产品的影响，以及验证状态和申报的工艺的影响)

| 3.3 Event repetitiveness review/事件重复性回顾 |  |  |
| --- | --- | --- |
| ☒ No repeat event in the past 12 month/ 过去十二个月内(from 10/16/2025 to 10/16/2026 )无同类事件发生 ☐This is a repeat event in the past 12 months/这是过去十二个月内(from N/A to N/A )重复发生的事件： 1) For the related event in the past, see/对于过去相关事件见: Deviation No./偏差编号： N/A 2) Confirm if the CAPAs are executed properly/确认之前的措施是否被正确执行： ☐ Yes/是 ☐ No/否 (a) If yes, proceed with this investigation for new action plan/如果是，继续该调查制定新行动项。 (b) If no, need the explanation, and whether new action shall be planned/如否，需要进行说明，以及是否需要执行新的行动项： N/A |  |  |
| Investigated by/Date/调查人/日期： |  |  |
| QA Confirm the Deviation Category/ QA确认偏差类别 | 1) Ranked as ☐ Critical ☐ Major ☐ Minor deviation is appropriate in page 1 of this deviation report. /本偏差报告第一页类别划分为☐关键 ☐主要☐次要偏差是合理的 2) The ranking in page 1 of this deviation report is impropriate per the later investigation, should be ranked as ☐ Critical ☐ Major ☐ Minor deviation, rationale: /根据后续的调查，本偏差文件第一页中偏差类别划分不合理，应该划分为☐关键 ☐主要 ☐次要偏差，理由是： |  |
|  | QA Approver /Date/ QA审批人/日期： |  |

| 4. Corrective and Preventative Action/纠正和预防措施 |  |  |
| --- | --- | --- |
| Note: Related Dept. personnel shall be involved to determine the corrective and preventative actions to ensure the effectiveness of CAPA. /注：纠正和预防措施应由涉及该偏差的相关部门人员提供以确保其有效。 |  |  |
| 4.1 Short-term action/短期措施 Short-term action shall be closed before 1st batch product release(if impact product release), or within 30 calendar days(if no impact on product release)./短期措施需要在首次产品放行前关闭(若影响产品放行)或者在30个自然日关闭(若不影响到产品放行) |  |  |
| Short-term corrective action/ 短期纠正措施 | Responsible person/落实人 | Due date / 到期时间 |
| - Clean bottom of bolt in the discharging port for synthesizer TJ4S-1213-Sy15, completed on 1017/2026, refer to Deviation-TJ420260226.01 Attachment 05; - Replace the package of sampler TJ4S-1213-Sy16-Sam01 with clean bag and remove the “cleaned” status label (will label as “Cleaned” after inspection), completed on 1017/2026, refer to Deviation-TJ420260226.01 Attachment 06; - Complete the final cleaning and release for TJ4S-1213-Sy15 and TJ4S-1213-Sy16, including sampler; | Yanlong Li | 1.Completed on 10/17/2026 2.Completed on 10/17/2026 3.10/21/2026 |
| (Note: please describe document no.. version no. and title if current doc. need to be updated./注：如需修订现行版本文件，请描述文件编号、版本号及标题) |  |  |
| Short-term Preventative action/ 短期预防措施 | Responsible person/落实人 | Due date / 到期时间 |
| Train all production person responsible for equipment cleaning and all QA for cleanliness inspection about Case 1, that the rear side of each bolt need to be cleaned and inspected. | Yanlong Li | 10/21/2026 |
| (Note: please describe document no.. version no. and title if current doc. need to be updated./注：如需修订现行版本文件，请描述文件编号、版本号及标题) |  |  |
| 4.2 Long-term action/长期措施 (Long term action shall not impact product release/长期措施不应影响产品放行) |  |  |
| Long-term corrective action/ 长期纠正措施 | Responsible person/落实人 | Due date / 到期时间 |
| Perform an extended check to all of peptide synthesizer in building 12, focusing on bottom of bolt in the discharging port. | Yanlong Li | 10/21/2026 |
| (Note: please describe document no.. version no. and title if current doc. need to be updated./注：如需修订现行版本文件，请描述文件编号、版本号及标题) |  |  |
| Long-term preventative action/ 长期预防措施 | Responsible person/落实人 | Due date / 到期时间 |
| - Updating RA-TJ42026048.01 <Risk Assessment for Cleaning of Peptide Solid-Phase Synthesizers at the TJ4 Site> to include bottom of bolt in the discharging port as the worst case cleaning points; - Updating cleaning record QF03376 <Cleaning and checking record - Solid Phase Peptide Synthesizer> to include the detailed check requirement for bottom of bolt in the discharging port; - Updating SOP DS-GMP-003 <Equipment cleaning management procedure for GMP production (DS)> give a more detailed instruction on the equipment dedicated accessory status label change; - Updating SOP DS-MM-015 <Processing aid materials management procedure (DS)> give a more detailed instruction on the gloves exchange during daily operation, including process operation and non-process operation. | 1.Xiongxiong 2.Anyao Tang 3.Xiongxiong 4.Xiongxiong | 1.10/25/2026 2.10/25/2026 3.10/25/2026 4.10/25/2026 |
| (Note: please describe document no.. version no. and title if current doc. need to be updated./注：如需修订现行版本文件，请描述文件编号、版本号及标题) |  |  |

| # 5. Revision history/版本历史 |  |  |
| --- | --- | --- |
| Version/ 版本号 | Approval date/ 审批日期 | Description of revision/修订描述 (including reason, background, content etc. /包括修订原因，背景，内容等） |
| V01 | Date of final signature on the deviation approval page | The first Version |

| 6. Deviation approval/偏差审批 |  |
| --- | --- |
| Authority ①/审批① |  |
| Project Leader/Date/ 项目主管/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |
| Dept. (except R&D) Reviewer/Date/ 部门（除研发）复核人/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |
| Authority ②/审批② |  |
| R&D senior director/Date/ 研发部高级主管/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |
| Dept. (except R&D) Approver/Date/ 部门（除研发）审批人/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |
| Authority ③/审批③ |  |
| Production director /Date/ 生产主管/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |
| Customer approval document attached/附有客户审批文件： ☐ Yes/是 ☐ No/否 ☐ N/A Further investigation should be performed/需要继续进行进一步调查： ☐ Yes/是 ☐ No/否 Further investigation includes/进一步调查包括： |  |
| QA Final Approver/Date/ QA最终审批人/日期 | ☐ Irrelevant/不涉及 ☐ Agree/同意 ☐ Disagree/不同意 （Signature/date/签字/日期） |

| 7. Deviation close-out/偏差的关闭 |  |  |  |
| --- | --- | --- | --- |
| 7.1 Technical close/技术关闭 (Short-term CAPA followup /短期CAPA跟踪) All short-term CAPA and “Task Prior to 1st impacted batch release” in the “section 3.1 Deviation impact on quality management system” shall be closed before the 1st release. /短期CAPA全部关闭、“3.1章节偏差对质量管理体系的影响”中“受影响第一批放行前需完成工作”全部完成，方可进行首次分发。 |  |  |  |
| Affected batch in the 1st release/首次分发受影响批次： |  |  |  |
| No. 序号 | Short-term CAPA/ 短期CAPA | Task completion date/ 工作完成日期 | Supporting Doc No./ 支持性文件编号 |
| No. 序号 | Code for “Task Prior to 1st impacted batch release” in the “section 3.1 / 3.1章节“受影响第一批放行前需完成工作”代号 | Task completion date/ 工作完成日期 | Supporting Doc No./ 支持性文件编号 |
| QA Reviewer/Date/QA复核人/日期： QA Approver /Date/QA审批人/日期： |  |  |  |

| 7.2 Final close/最终关闭（Long-term CAPA followup/长期CAPA跟踪） For the long-term CAPA and "Post release Task” in the “section 3.1, the open actions can be transferred into monthly CAPA tracking system to follow up until closed. /长期CAPA及3.1章节中放行后需要完成的工作，请在下表中记录最新进展，对于未关闭的工作，转入月CAPA跟踪系统中进行跟踪。 |  |  |  |  |
| --- | --- | --- | --- | --- |
| No. 序号 | Long-term CAPA/ 长期CAPA |  | Task completion date/ 工作完成日期 | Supporting Doc No./ 支持性文件编号 |
| No. 序号 | Code for "Post release Task” in the “section 3.1/ 3.1章节中放行后需要完成工作代号 |  | Task completion date/ 工作完成日期 | Supporting Doc No./ 支持性文件编号 |
| Initial CAPA tracking report No./所有未关闭项首次加入CAPA跟踪报告的编号： |  |  |  |  |
| QA Reviewer/Date/QA复核人/日期： QA Approver /Date/QA审批人/日期： |  |  |  |  |
| Confirmation of final closure /最终关闭确认 |  |  |  |  |
| Unclosed task in CAPA tracking / 加入CAPA跟踪报告的未关闭项 |  | Task completion date/ 工作完成日期 | Supporting Doc No./ 支持性文件编号 | Confirmed by /Date/ QA确认人/日期 |

Note/注：

a. The numbering rule for process deviation document is Deviation - Project No. + Step No. + Sequential No..two-digital version no. , e.g., the first approval version for the second process deviation of project CPo121234-01 step 02 is Deviation -CPo121234-010202.01 . In particular, the number rule for drug product business unit is Deviation - site code +PD+project number+3-digit serial no. two-digital version no., such as Deviation –TJ2PDPM301234-01001.01. /生产工艺偏差编号原则：Deviation-项目编号＋步骤号＋序号.两位版本号，例如： CPo121234-01项目第二步发生的第二个生产工艺偏差审批的第一个版本编号为Deviation -CPo121234-010202.01。特别的，对于制剂业务模块为Deviation –厂区代码+PD+项目编号+3位序号.两位版本号，如Deviation–TJ2PDPM301234-01001.01。

b. The numbering rule for other Deviation document is Deviation -Site Code (identification code shall be added for the business unit other than DS) + Year+ Sequential No. two- digital version no., e.g., Deviation -TJ220130001.01, Deviation -TJ2PD2013001.01. /生产工艺以外的偏差编号原则：Deviation -厂区代码(原料药以外的业务模块增加识别码)＋年份＋序号.两位版本号，例如：Deviation -TJ220130001.01，Deviation -TJ2PD2013001.01。